## Intro & TOC
Sea-Bird Electronics Moored Instrument Battery Budget Calculator v1.0 |  |  | 2004-08-20
Introduction
This power budget calculator is published as an aid in determining the number of samples for a specified sample duration or rate,
based on Sea-Bird's conservative estimate of battery capacity and an instrument's memory capacity. Our battery capacity estimates are based
on manufacturer's published capacity, and adjusted (de-rated) for cold temperature and cut-off voltage characteristics of particular instruments.
Values for current consumption used in these calculators are taken from our product manuals and third party manufacturer's data sheets.
We have measured all of the performance parameters of the instruments and their respective configurations of current SBE production instrumentation,
and we believe these values are conservative. Our best effort has been made to provide a useful deployemnt planning tool, while minimizing the
risk of data loss due to overly optimistic battery capacity assumptions.
Note: For SBE 26 and 26plus Wave and Tide Recorder - use the power budget calculation module incorporated into SEASOFT for WAVES
GENERAL DISCLAIMER
Due to variations in components, instrumentation, production revisions, battery age or shelf life and general factors beyond our control,
Sea-Bird Electronics offers this calculator FOR USE AT THE END USER'S OWN RISK. Instrument battery life (like automobile gasoline mileage)
may vary in any given environment.
Notes
Product Manual examples and values have been used to confirm these calculations and the Manual revision is noted.
Some 'default' values in this calculator may vary from published values - we will attempt to note differences in the specific configuration "Notes:"
Updates, corrections, features, etc.
Please email suggestions, etc. to Brian via: bmccully@seabird.com
Jump link: | Instrument and configuration | note: the various worksheet instrument configurations may also be accessed via
(click on the > below) |  | clicking on the tabs below the sheet
> | 16Plus CT no pump | the tab names may also be edited (renamed, copied, moved, etc.) by
> | 16Plus CTD no pump | right-clicking on the tab
> | 16Plus CT, pumped
> | 16Plus CTD, pumped
> | 16PIM CT no pump
> | 16PIM CTD no pump
> | 16PIM CT, pumped
> | 16PIM CTD, pumped
> | 16Plus RS-485 CT no pump
> | 16Plus RS-485 CTD no pump
> | 16Plus RS-485 CT pumped
> | 16Plus RS-485 CTD pumped
> | SBE 37IM MicroCAT CT Recorder
> | SBE 37IM MicroCAT CTD Recorder
> | SBE 37IMP MicroCAT CT Recorder
> | SBE 37IMP MicroCAT CTD Recorder
> | SBE 37SM MicroCAT CT Recorder, internal power
> | SBE 37SM MicroCAT CT Recorder, external power
> | SBE 37SM MicroCAT CTD Recorder, internal power
> | SBE 37SM MicroCAT CTD Recorder, external power
> | SBE 37SM MicroCAT CT Recorder, RS-485
> | SBE 37SM MicroCAT CTD Recorder, RS-485
> | SBE 37SMP MicroCAT CT Recorder, internal power
> | SBE 37SMP MicroCAT CT Recorder, external power
> | SBE 37SMP MicroCAT CTD Recorder, internal power
> | SBE 37SMP MicroCAT CTD Recorder, external power
> | SBE 37SMP MicroCAT CT Recorder, RS-485
> | SBE 37SMP MicroCAT CTD Recorder, RS-485
> | SBE 39 Temperature (optional P) Recorder - interval
> | SBE 39 Temperature (optional P) Recorder - continuous
> | SBE 44 Underwater Inductive Modem
> | SBE 16 SEACAT CT Recorder, no pump
> | SBE 16 SEACAT CTD Recorder, no pump
> | SBE 16 SEACAT CT Recorder, with pump
> | SBE 16 SEACAT CTD Recorder, with pump
> | SBE 16 SEACAT CT & latched DO Recorder
> | SBE 16 SEACAT CTD & latched DO Recorder

## How To...
General Instructions
This calculator is based upon the Specifications, Battery Endurance and Sample Timing information and examples in the product manuals
Green fields | are user-entry fields for the variety of deployment schemes.
Blue fields | are calculated values used for programming the CTD or for general information
Peach | fields are used for gauging available memory or battery life
This calculator file is intended to be "saved as" and tailored for your application - i.e. it is not a read only file.
If you wish to use areas outside of the user entry fields (for example, for scratchpad user notes), change the Tools>Protection>Unprotect Sheet status
Note: some fields to the right of the visible spreadsheet area are used for data or calculations - do not overwrite or the spreadsheet may not function properly
The Target Duration function (not found in the product manual) can be used as a general 'gas gauge' determination
This function can also be used for optimized sampling - i.e. setting up the fastest sample rate for a target deployment span.
The Water temperature function applies for calculations using Alkaline batteries, for which capacity decreases with colder water temperatures.

## 16Plus CT no pump
SBE 16Plus Power Budget Calculations |  | CT, no pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
[6 empty] | Target duration = | 9 | months (~30 days/m)
Sampling duration (default 2.2 seconds) |  |  |  |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  |  | capacity (AH) | days or | months
0.05 | 1 | 0.11 | 12 | 1.428 | 12.2 | 1281.5126 | 42.7171 | 369075.6303 | 21.07%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
| 2.2 |  |  | (+quiescent+auxes) | use 10.5 amp-hours for pump or aux sensors
| sampling time |  | 5 | (min) interval
|  |  | 300 | (sec) interval
Delay before sampling | 0 |  | (sampleinterval = x)
current draw (mA) |  |  |  |  | Memory |  | bytes used | Available
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | T&C | 6 | 6 | Samples
Aux 2 | 0 | then enter a value = 45 |  |  | Aux 1 | 2 | 0 | 800000
Aux 3 | 0 |  |  |  | Aux 2 | 2 | 0
Aux 4 | 0 |  |  |  | Aux 3 | 2 | 0
Sum of Auxes | 0 | mA |  |  | Aux 4 | 2 | 0
[5 empty] | SBE 38 or 50 | 3 | 0
Quiescent consumption in AH |  |  |  |  | Date/Time | 4 | 4
3e-05 | 3600 | 0.108 |  |  | GTD 1 | 7 | 0
amps | sec/hr | amp-sec per hr |  |  | GTD 2 | 7 | 0
[5 empty] | GTD 2 | 7 | 0 |  | % used
using Power endurance examples SBE 16Plus manual v11 |  |  |  |  | Sum |  | 10 |  | 9.72%
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3))

## 16Plus CTD no pump
SBE 16Plus Power Budget Calculations |  | CTD, no pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
[6 empty] | Target duration = | 9 | months (~30 days/m) |  | 1
Sampling duration (default 2.5 seconds) |  |  |  |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  |  | capacity (AH) | days or | months
0.065 | 1 | 0.208 | 20 | 4.268 | 12.2 | 428.7723 | 14.2924 | 205810.6842 | 62.97%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
Parosintegration= | 1 | (sec) |  | (+quiescent+auxes) | use 10.5 amp-hours for pump or aux sensors
| range is 1 to 6000 |  | 3 | (min) interval
| 3.2 |  | 180 | (sec) interval
| sampling time |  | (sampleinterval = x)
Delay before sampling | 0 |  |  |  | Memory |  | bytes used | Available
current draw (mA) |  |  |  |  | T&C | 6 | 6 | Samples
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | Druck | 5 | 0 | 500000
Aux 2 | 0 | then enter a value = 45 |  |  | Digiquartz | 6 | 6
Aux 3 | 0 |  |  |  | Aux 1 | 2 | 0
Aux 4 | 0 |  |  |  | Aux 2 | 2 | 0
Sum of Auxes | 0 | mA |  |  | Aux 3 | 2 | 0
[5 empty] | Aux 4 | 2 | 0
Quiescent consumption in AH |  |  |  |  | SBE 38 or 50 | 3 | 0
3e-05 | 3600 | 0.108 |  |  | Date/Time | 4 | 4
amps | sec/hr | amp-sec per hr |  |  | GTD 1 | 7 | 0
[5 empty] | GTD 2 | 7 | 0
[5 empty] | GTD 2 | 7 | 0 |  | % used
[5 empty] | Sum |  | 16 |  | 25.92%
Notes:
Deployments of alkaline batteries longer than 2 years are not advised
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3))

## 16Plus CT, pumped
SBE 16Plus Power Budget Calculations |  | CT, with pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
[6 empty] | Target duration = | 9 | months (~30 days/m)
Sampling duration |  |  |  |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  | Total | capacity (AH) | days or | months
0.05 | 1 | 0.135 | 12 | 2.328 | 10.5 | 676.5464 | 22.5515 | 194845.3608 | 39.91%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
ncycle sum (sec) | 2.2 |  | 5 | (min) interval | use 10.5 amp-hours for pump or aux sensors
sampling time | 2.7 | sec | 300 | (sec) interval
(sampleinterval = x)
5M (or 5T) current | (default is 0.5 sec)
0.1 | mode 1 | 0.05 |  |  | Memory |  | bytes used | Available |  |  | 0.1
0.1 (or 0.15) amps | **pump mode | amp-sec per hr |  |  | T&C | 6 | 6 | Samples |  |  | 0.15
[5 empty] | Aux 1 | 2 | 0 | 800000
sampling delay (s) | 0 | 0.5 | pump time (sec) |  | Aux 2 | 2 | 0
| delaybeforesampling=x |  |  |  | Aux 3 | 2 | 0
current draw (mA) |  |  |  |  | Aux 4 | 2 | 0 |  |  | mode 1 | 0.5 | mode 1 | 0.5
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | SBE 38 or 50 | 3 | 0 |  |  | mode 2 | 2.2
Aux 2 | 0 | then enter a value = 45 |  |  | Date/Time | 4 | 4
Aux 3 | 0 |  |  |  | GTD 1 | 7 | 0
Aux 4 | 0 |  |  |  | GTD 2 | 7 | 0
Sum of Auxes | 0 | mA |  |  | GTD 2 | 7 | 0 |  | % used
[5 empty] | Sum |  | 10 |  | 9.72%
Quiescent consumption in AH
3e-05 | 3600 | 0.108
amps | sec/hr | amp-sec per hr
using Power endurance examples SBE 16Plus manual v11
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3))
** Pump mode 1 does not incorporate the 'delay before sampling' values - this operation would rarely if ever be used

## 16Plus&19Plus CTD, pumped
SBE 16Plus/19Plus Power Calculations |  | CTD, with pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
Note: For SBE 19Plus moored mode calculation, set Parosintegration to zero value. | [5 empty] | Target duration = | 6 | months (~30 days/m) |  |  | 0.3
Sampling duration |  |  |  |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  | Total | capacity (AH) | days or | months
0.065 | 1 | 0.195 | 20 | 5.508 | 10.5 | 285.9477 | 9.5316 | 137254.902 | 62.95%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
Parosintegration= | 0 | (sec) |  |  | use 10.5 amp-hours for pump or aux sensors
| range is 1 to 6000 |  | 3 | (min) interval
ncycle sum (sec) | 2.5 |  | 180 | (sec) interval
sampling time | 3 | sec | (sampleinterval = x)
[5 empty] | Memory |  | bytes used | Available |  |  | 0.1
5M (or 5T) current | (default is 0.5 sec) |  |  |  | T&C | 6 | 6 | Samples |  |  | 0.15
0.15 | mode 1 | 0.075 |  |  | Druck | 5 | 5 | 533333.3333
0.1 (or 0.15) amps | **pump mode | amp-sec per hr |  |  | Digiquartz | 6 | 0
[5 empty] | Aux 1 | 2 | 0
sampling delay (s) | 20 | 0.5 | pump time (sec) |  | Aux 2 | 2 | 0 |  |  | mode 1 | 0.5 | mode 1 | 0.5
| delaybeforesampling=x |  |  |  | Aux 3 | 2 | 0 |  |  | mode 2 | 22.5
current draw (mA) |  |  |  |  | Aux 4 | 2 | 0
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | SBE 38 or 50 | 3 | 0
Aux 2 | 0 | then enter a value = 45 |  |  | Date/Time | 4 | 4
Aux 3 | 0 |  |  |  | GTD 1 | 7 | 0
Aux 4 | 0 |  |  |  | GTD 2 | 7 | 0
Sum of Auxes | 0 | mA |  |  | GTD 2 | 7 | 0 |  | % used
[5 empty] | Sum |  | 15 |  | 16.2%
Quiescent consumption in AH
3e-05 | 3600 | 0.108
amps | sec/hr | amp-sec per hr
using Power endurance examples SBE 16Plus manual v11
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3))
** Pump mode 1 does not incorporate the 'delay before sampling' values - this operation would rarely if ever be used

## 16PIM CT no pump
SBE 16Plus-IM Power Calculations |  | CT, no pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
Commun. Current | 10 | 0.025 | 1
| # of moored instruments | amp-sec per hr | requests/hr |  |  | Target duration = | 9 | months (~30 days/m)
Sampling duration (default 2.2 seconds) |  |  |  |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  |  | capacity (AH) | days or | months
0.05 | 1 | 0.11 | 10 | 1.611 | 12.2 | 1135.9404 | 37.8647 | 272625.6983 | 23.77%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
| 2.2 |  |  | (+quiescent+auxes) | use 10.5 amp-hours for pump or aux sensors
| sampling time |  | 6 | (min) interval
|  |  | 360 | (sec) interval
Delay before sampling | 0 |  | (sampleinterval = x)
current draw (mA) |  |  |  |  | Memory |  | bytes used | Available
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | T&C | 6 | 6 | Samples
Aux 2 | 0 | then enter a value = 45 |  |  | Aux 1 | 2 | 0 | 800000
Aux 3 | 0 |  |  |  | Aux 2 | 2 | 0
Aux 4 | 0 |  |  |  | Aux 3 | 2 | 0
Sum of Auxes | 0 | mA |  |  | Aux 4 | 2 | 0
[5 empty] | SBE 38 or 50 | 3 | 0
Quiescent consumption in AH |  |  |  |  | Date/Time | 4 | 4
0.000135 | 3600 | 0.486 |  |  | GTD 1 | 7 | 0
amps | sec/hr | amp-sec per hr |  |  | GTD 2 | 7 | 0
[5 empty] | GTD 2 | 7 | 0 |  | % used
using Power endurance examples SBE 16Plus manual v11 |  |  |  |  | Sum |  | 10 |  | 8.1%
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3))

## 16PIM CTD no pump
SBE 16Plus-IM Power Calculations |  | CTD, no pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
Commun. Current | 10 | 0.025 | 1 |  |  | Target duration = | 9 | months (~30 days/m) |  | 0.3
| # of moored instruments | amp-sec per hr | requests/hr
Sampling duration (default 2.5 seconds) |  |  |  |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  |  | capacity (AH) | days or | months
0.065 | 1 | 0.1625 | 40 | 7.011 | 12.2 | 261.0184 | 8.7006 | 250577.6637 | 103.4%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
Druck (0) or Digi (1) | 0 |  |  | (+quiescent+auxes) | use 10.5 amp-hours for pump or aux sensors
| 2.5 |  | 1.5 | (min) interval
| sampling time |  | 90 | (sec) interval
Delay before sampling | 0 |  | (sampleinterval = x)
current draw (mA) |  |  |  |  | Memory |  | bytes used | Available
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | T&C | 6 | 6 | Samples
Aux 2 | 0 | then enter a value = 45 |  |  | Druck | 5 | 0 | 500000
Aux 3 | 0 |  |  |  | Digiquartz | 6 | 6
Aux 4 | 0 |  |  |  | Aux 1 | 2 | 0
Sum of Auxes | 0 | mA |  |  | Aux 2 | 2 | 0
[5 empty] | Aux 3 | 2 | 0
Quiescent consumption in AH |  |  |  |  | Aux 4 | 2 | 0
0.000135 | 3600 | 0.486 |  |  | SBE 38 or 50 | 3 | 0
amps | sec/hr | amp-sec per hr |  |  | Date/Time | 4 | 4
[5 empty] | GTD 1 | 7 | 0
[5 empty] | GTD 2 | 7 | 0
[5 empty] | GTD 2 | 7 | 0 |  | % used
[5 empty] | Sum |  | 16 |  | 51.84%
Notes:
Deployments of alkaline batteries longer than 2 years are not advised
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3))

## 16PIM CT, pumped
SBE 16Plus-IM Power Calculations |  | CT, with pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
Commun. Current | 10 | 0.025 | 1 |  |  | Target duration = | 9 | months (~30 days/m)
| # of moored instruments | amp-sec per hr | requests/hr
Sampling duration |  |  |  |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  | Total | capacity (AH) | days or | months
0.05 | 1 | 0.135 | 6 | 1.621 | 10.5 | 971.6225 | 32.3874 | 139913.6336 | 27.79%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
ncycle sum (sec) | 2.2 |  | 10 | (min) interval | use 10.5 amp-hours for pump or aux sensors
sampling time | 2.7 | sec | 600 | (sec) interval
(sampleinterval = x)
5M (or 5T) current | (default is 0.5 sec)
0.1 | mode 1 | 0.05 |  |  | Memory |  | bytes used | Available |  |  | 0.1
0.1 (or 0.15) amps | **pump mode | amp-sec per hr |  |  | T&C | 6 | 6 | Samples |  |  | 0.15
[5 empty] | Aux 1 | 2 | 0 | 800000
sampling delay (s) | 0 | 0.5 | pump time (sec) |  | Aux 2 | 2 | 0
| #iidelaybeforesampling=x |  |  |  | Aux 3 | 2 | 0
current draw (mA) |  |  |  |  | Aux 4 | 2 | 0 |  |  | mode 1 | 0.5 | mode 1 | 0.5
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | SBE 38 or 50 | 3 | 0 |  |  | mode 2 | 2.2
Aux 2 | 0 | then enter a value = 45 |  |  | Date/Time | 4 | 4
Aux 3 | 0 |  |  |  | GTD 1 | 7 | 0
Aux 4 | 0 |  |  |  | GTD 2 | 7 | 0
Sum of Auxes | 0 | mA |  |  | GTD 2 | 7 | 0 |  | % used
[5 empty] | Sum |  | 10 |  | 4.86%
Quiescent consumption in AH
0.000135 | 3600 | 0.486
amps | sec/hr | amp-sec per hr
using Power endurance examples SBE 16Plus manual v11
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3))
** Pump mode 1 does not incorporate the 'delay before sampling' values - this operation would rarely if ever be used

## 16PIM CTD, pumped
SBE 16Plus-IM Power Calculations |  | CTD, with pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
Commun. Current | 10 | 0.025 | 1 |  |  | Target duration = | 1 | months (~30 days/m) |  |  | 0
| # of moored instruments | amp-sec per hr | requests/hr
Sampling duration |  |  |  |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  | Total | capacity (AH) | days or | months
0.065 | 1 | 2.838 | 6 | 33.019 | 10.5 | 47.6998 | 1.59 | 6868.7725 | 62.89%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
Druck (0) or Digi (1) | 1 |  |  |  | use 10.5 amp-hours for pump or aux sensors
ncycle sum (sec) | 2.2 |  | 10 | (min) interval
sampling time | 17.2 | sec | 600 | (sec) interval
(sampleinterval = x)
[5 empty] | Memory |  | bytes used | Available |  |  | 0.1
5M (or 5T) current | (default is 0.5 sec) |  |  |  | T&C | 6 | 6 | Samples |  |  | 0.15
0.15 | mode 2 | 2.58 |  |  | Druck | 5 | 0 | 333333.3333
0.1 (or 0.15) amps | **pump mode | amp-sec per hr |  |  | Digiquartz | 6 | 6
[5 empty] | Aux 1 | 2 | 2
sampling delay (s) | 15 | 17.2 | pump time (sec) |  | Aux 2 | 2 | 2 |  |  | mode 1 | 0.5 | mode 1 | 0.5
| #iidelaybeforesampling=x |  |  |  | Aux 3 | 2 | 2 |  |  | mode 2 | 17.2
current draw (mA) |  |  |  |  | Aux 4 | 2 | 2
Aux 1 | 100 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | SBE 38 or 50 | 3 | 0
Aux 2 | 0 | then enter a value = 45 |  |  | Date/Time | 4 | 4
Aux 3 | 0 |  |  |  | GTD 1 | 7 | 0
Aux 4 | 0 |  |  |  | GTD 2 | 7 | 0
Sum of Auxes | 100 | mA |  |  | GTD 2 | 7 | 0 |  | % used
[5 empty] | Sum |  | 24 |  | 1.296%
Quiescent consumption in AH
0.000135 | 3600 | 0.486
amps | sec/hr | amp-sec per hr
using Power endurance examples SBE 16Plus manual v11
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3)) and only even increments are used with Digiquartz (2,4,6, etc.)
** Pump mode 1 does not incorporate the 'delay before sampling' values - this operation would rarely if ever be used

## 16Plus RS-485 CT no pump
SBE 16Plus RS-485 Power Calculations |  | CT, no pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
Commun. Current | 10 | 0.0065 | 1 |  |  | Target duration = | 9 | months (~30 days/m)
| # of moored instruments | amp-sec per hr | requests/hr
Sampling duration (default 2.2 seconds) |  |  |  |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  |  | capacity (AH) | days or | months
0.05 | 1 | 0.11 | 6 | 0.7745 | 12.2 | 2362.8147 | 78.7605 | 340245.3196 | 11.43%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
| 2.2 |  |  | (+quiescent+auxes) | use 10.5 amp-hours for pump or aux sensors
| sampling time |  | 10 | (min) interval
|  |  | 600 | (sec) interval
Delay before sampling | 0 |  | (sampleinterval = x)
current draw (mA) |  |  |  |  | Memory |  | bytes used | Available
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | T&C | 6 | 6 | Samples
Aux 2 | 0 | then enter a value = 45 |  |  | Aux 1 | 2 | 0 | 800000
Aux 3 | 0 |  |  |  | Aux 2 | 2 | 0
Aux 4 | 0 |  |  |  | Aux 3 | 2 | 0
Sum of Auxes | 0 | mA |  |  | Aux 4 | 2 | 0
[5 empty] | SBE 38 or 50 | 3 | 0
Quiescent consumption in AH |  |  |  |  | Date/Time | 4 | 4
3e-05 | 3600 | 0.108 |  |  | GTD 1 | 7 | 0
amps | sec/hr | amp-sec per hr |  |  | GTD 2 | 7 | 0
[5 empty] | GTD 2 | 7 | 0 |  | % used
using Power endurance examples SBE 16Plus manual v11 |  |  |  |  | Sum |  | 10 |  | 4.86%
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3))

## 16Plus RS-485 CTD no pump
SBE 16Plus RS-485 Power Calculations |  | CTD, no pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
Commun. Current | 10 | 0.0065 | 1 |  |  | Target duration = | 9 | months (~30 days/m) |  | 1
| # of moored instruments | amp-sec per hr | requests/hr
Sampling duration (default 2.5 seconds) |  |  | (default of 1 sample per hour) |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  |  | capacity (AH) | days or | months
0.065 | 1 | 0.208 | 20 | 4.2745 | 12.2 | 428.1202 | 14.2707 | 205497.719 | 63.07%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
Parosintegration= | 1 | (sec) |  | (+quiescent+auxes) | use 10.5 amp-hours for pump or aux sensors
| range is 1 to 6000 |  | 3 | (min) interval
| 3.2 |  | 180 | (sec) interval
| sampling time |  | (sampleinterval = x)
Delay before sampling | 0 |  |  |  | Memory |  | bytes used | Available
current draw (mA) |  |  |  |  | T&C | 6 | 6 | Samples
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | Druck | 5 | 0 | 500000
Aux 2 | 0 | then enter a value = 45 |  |  | Digiquartz | 6 | 6
Aux 3 | 0 |  |  |  | Aux 1 | 2 | 0
Aux 4 | 0 |  |  |  | Aux 2 | 2 | 0
Sum of Auxes | 0 | mA |  |  | Aux 3 | 2 | 0
[5 empty] | Aux 4 | 2 | 0
Quiescent consumption in AH |  |  |  |  | SBE 38 or 50 | 3 | 0
3e-05 | 3600 | 0.108 |  |  | Date/Time | 4 | 4
amps | sec/hr | amp-sec per hr |  |  | GTD 1 | 7 | 0
[5 empty] | GTD 2 | 7 | 0
[5 empty] | GTD 2 | 7 | 0 |  | % used
[5 empty] | Sum |  | 16 |  | 25.92%
Notes:
Deployments of alkaline batteries longer than 2 years are not advised
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3))

## 16Plus RS-485 CT, pumped
SBE 16Plus RS-485 Power Calculations |  | CT, with pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
Commun. Current | 10 | 0.0065 | 1 |  |  | Target duration = | 9 | months (~30 days/m)
| # of moored instruments | amp-sec per hr | requests/hr
Sampling duration |  |  |  |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  | Total | capacity (AH) | days or | months
0.05 | 1 | 0.135 | 6 | 1.2245 | 10.5 | 1286.2393 | 42.8746 | 185218.4565 | 20.99%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
ncycle sum (sec) | 2.2 |  | 10 | (min) interval | use 10.5 amp-hours for pump or aux sensors
sampling time | 2.7 | sec | 600 | (sec) interval
(sampleinterval = x)
5M (or 5T) current | (default is 0.5 sec)
0.1 | mode 1 | 0.05 |  |  | Memory |  | bytes used | Available |  |  | 0.1
0.1 (or 0.15) amps | **pump mode | amp-sec per hr |  |  | T&C | 6 | 6 | Samples |  |  | 0.15
[5 empty] | Aux 1 | 2 | 0 | 800000
sampling delay (s) | 0 | 0.5 | pump time (sec) |  | Aux 2 | 2 | 0
| delaybeforesampling=x |  |  |  | Aux 3 | 2 | 0
current draw (mA) |  |  |  |  | Aux 4 | 2 | 0 |  |  | mode 1 | 0.5 | mode 1 | 0.5
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | SBE 38 or 50 | 3 | 0 |  |  | mode 2 | 2.2
Aux 2 | 0 | then enter a value = 45 |  |  | Date/Time | 4 | 4
Aux 3 | 0 |  |  |  | GTD 1 | 7 | 0
Aux 4 | 0 |  |  |  | GTD 2 | 7 | 0
Sum of Auxes | 0 | mA |  |  | GTD 2 | 7 | 0 |  | % used
[5 empty] | Sum |  | 10 |  | 4.86%
Quiescent consumption in AH
3e-05 | 3600 | 0.108
amps | sec/hr | amp-sec per hr
using Power endurance examples SBE 16Plus manual v11
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3))
** Pump mode 1 does not incorporate the 'delay before sampling' values - this operation would rarely if ever be used

## 16Plus RS-485 CTD, pumped
SBE 16Plus RS-485 Power Calculations |  | CTD, with pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
Commun. Current | 10 | 0.0065 | 1 |  |  | Target duration = | 1 | months (~30 days/m) |  |  | 3
| # of moored instruments | amp-sec per hr | requests/hr
Sampling duration |  |  |  |  | Battery | deployment |  | # samples | % used
| (enter x number of Ncycles) |  |  | Total | capacity (AH) | days or | months
0.065 | 4 | 3.4567 | 12 | 79.3055 | 10.5 | 19.8599 | 0.662 | 5719.6537 | 151.1%
amps | *range is 1 to 255 | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
Parosintegration= | 3 | (sec) |  |  | use 10.5 amp-hours for pump or aux sensors
| range is 1 to 6000 |  | 5 | (min) interval
ncycle sum (sec) | 5.95 |  | 300 | (sec) interval
sampling time | 20.95 | sec | (sampleinterval = x)
[5 empty] | Memory |  | bytes used | Available |  |  | 0.1
5M (or 5T) current | (default is 0.5 sec) |  |  |  | T&C | 6 | 6 | Samples |  |  | 0.15
0.15 | mode 2 | 3.1425 |  |  | Druck | 5 | 0 | 333333.3333
0.1 (or 0.15) amps | **pump mode | amp-sec per hr |  |  | Digiquartz | 6 | 6
[5 empty] | Aux 1 | 2 | 2
sampling delay (s) | 15 | 20.95 | pump time (sec) |  | Aux 2 | 2 | 2 |  |  | mode 1 | 0.5 | mode 1 | 0.5
| delaybeforesampling=x |  |  |  | Aux 3 | 2 | 2 |  |  | mode 2 | 20.95
current draw (mA) |  |  |  |  | Aux 4 | 2 | 2
Aux 1 | 100 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | SBE 38 or 50 | 3 | 0
Aux 2 | 0 | then enter a value = 45 |  |  | Date/Time | 4 | 4
Aux 3 | 0 |  |  |  | GTD 1 | 7 | 0
Aux 4 | 0 |  |  |  | GTD 2 | 7 | 0
Sum of Auxes | 100 | mA |  |  | GTD 2 | 7 | 0 |  | % used
[5 empty] | Sum |  | 24 |  | 2.592%
Quiescent consumption in AH
3e-05 | 3600 | 0.108
amps | sec/hr | amp-sec per hr
using Power endurance examples SBE 16Plus manual v11
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
*Maximum number of Ncycles is the smaller of Ncycles = 1 to 255 or (4*(sampleinterval-3))
** Pump mode 1 does not incorporate the 'delay before sampling' values - this operation would rarely if ever be used

## 37IM CT
SBE 37IM Power Budget Calculations |  | CT, no pressure | (enter variables into green boxes)
(click here to -> Intro)
Number of Instruments on line
| 10 | 0.025 | 1 |  |  | Target duration = | 12 | months (~30 days/m)
|  | amp-sec per hr | requests/hr
[5 empty] | Battery | deployment |  | # samples | % used
Autonomous or Serial Line Sync Sampling |  |  |  |  | capacity (AH) | days or | months
0.03 | 2.2 | 0.066 | 18 | 1.573 | 5 | 476.7959 | 15.8932 | 205975.8423 | 75.5%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 3.3333 | (min) sampling interval
|  |  | 200 | (sec) sampling interval
(Interval =x)
[5 empty] | Battery | deployment |  | # samples | % used
Polled Sampling |  |  |  |  | capacity (AH) | days or | months
0.03 | 2.95 | 0.0885 | 18 | 1.978 | 5 | 379.1709 | 12.639 | 163801.82 | 94.94%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 3.3333 | (min) sampling interval
|  |  | 200 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  | Capacity (Kbytes) | % used | 0
0.0001 | 3600 | 0.36 |  |  | T&C | 5 |  | 2048 | Autonom. | 4
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 68.34%
[5 empty] | (no) Druck | 0 |  | Samples | Polled | 1024
Using Power endurance examples SBE 37IM manual v13, w/ 1.75 seconds added to polled sample |  |  |  |  | Per Sample | 9 |  | 227555.5556 | 68.34% | 2048
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37IM CTD
SBE 37IM Power Budget Calculations |  | CT with pressure | (enter variables into green boxes)
(click here to -> Intro)
Number of Instruments on line
| 10 | 0.025 | 1 |  |  | Target duration = | 12 | months (~30 days/m)
|  | amp-sec per hr | requests/hr
[5 empty] | Battery | deployment |  | # samples | % used
Autonomous or Serial Line Sync Sampling |  |  |  |  | capacity (AH) | days or | months
0.03 | 2.7 | 0.081 | 15 | 1.6 | 5 | 468.75 | 15.625 | 168750 | 76.8%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 4 | (min) sampling interval
|  |  | 240 | (sec) sampling interval
(Interval =x)
[5 empty] | Battery | deployment |  | # samples | % used
Polled Sampling |  |  |  |  | capacity (AH) | days or | months
0.03 | 3.35 | 0.1005 | 6 | 0.988 | 5 | 759.1093 | 25.3036 | 109311.7409 | 47.42%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  | Capacity (Kbytes) | % used | 0
0.0001 | 3600 | 0.36 |  |  | T&C | 5 |  | 2048 | Autonom. | 4
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 69.61%
[5 empty] | (no) Druck | 2 |  | Samples | Polled | 1024
Using Power endurance examples SBE 37IM manual v13, w/ 1.75 seconds added to polled sample |  |  |  |  | Per Sample | 11 |  | 186181.8182 | 27.84% | 2048
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37IMP CT
SBE 37IMP Power Budget Calculations |  | CT, no pressure | (enter variables into green boxes)
(click here to -> Intro)
Number of Instruments on line
| 10 | 0.025 | 1 |  |  | Target duration = | 12 | months (~30 days/m)
|  | amp-sec per hr | requests/hr
[5 empty] | Battery | deployment |  | # samples | % used
Autonomous or Serial Line Sync Sampling |  |  |  |  | capacity (AH) | days or | months
0.03 | 2.7 | 0.081 | 15 | 3.68 | 5 | 203.8043 | 6.7935 | 73369.5652 | 176.6%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 4 | (min) sampling interval
|  |  | 240 | (sec) sampling interval
(Interval =x)
[5 empty] | Battery | deployment |  | # samples | % used
Polled Sampling |  |  |  |  | capacity (AH) | days or | months
0.03 | 3.45 | 0.1035 | 10 | 2.85 | 5 | 263.1579 | 8.7719 | 63157.8947 | 136.8%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 6 | (min) sampling interval
|  |  | 360 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  |  | % used | 0
0.0001 | 3600 | 0.36 |  |  | T&C | 5 |  |  | Autonom. | 4
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 56.95%
Pump current consumption for 0.5 sec |  |  |  |  | (no) Druck | 0 |  | Samples | Polled | 1024
0.26 | 0.5 | 0.13 |  |  | Per Sample | 9 |  | 227555.5556 | 37.97% | 2048
amps | sampling duration (s) | amp-sec per hr
Using Power endurance examples SBE 37IMP manual v3, w/ 1.75 seconds added to polled sample
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37IMP CTD
SBE 37IMP Power Budget Calculations |  | CT with pressure | (enter variables into green boxes)
(click here to -> Intro)
Number of Instruments on line
| 10 | 0.025 | 1 |  |  | Target duration = | 12 | months (~30 days/m)
|  | amp-sec per hr | requests/hr
[5 empty] | Battery | deployment |  | # samples | % used
Autonomous or Serial Line Sync Sampling |  |  |  |  | capacity (AH) | days or | months
0.03 | 3.2 | 0.096 | 16 | 4.131 | 5 | 181.5541 | 6.0518 | 69716.7756 | 198.3%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 3.75 | (min) sampling interval
|  |  | 225 | (sec) sampling interval
(Interval =x)
[5 empty] | Battery | deployment |  | # samples | % used
Polled Sampling |  |  |  |  | capacity (AH) | days or | months
0.03 | 3.85 | 0.1155 | 12 | 3.461 | 5 | 216.7004 | 7.2233 | 62409.7082 | 166.1%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 5 | (min) sampling interval
|  |  | 300 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  |  | % used | 0
0.0001 | 3600 | 0.36 |  |  | T&C | 5 |  |  | Autonom. | 4
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 74.25%
Pump current consumption for 0.5 sec |  |  |  |  | (no) Druck | 2 |  | Samples | Polled
0.26 | 0.5 | 0.13 |  |  | Per Sample | 11 |  | 186181.8182 | 55.69%
amps | sampling duration (s) | amp-sec per hr
Using Power endurance examples SBE 37IMP manual v3, w/ 1.75 seconds added to polled sample
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SM CT intpower
SBE 37SM Power Budget Calculations |  | CT, no pressure, internal power
(enter variables into green boxes) |  | (click here to -> Intro) |  |  |  | Target duration = | 6 | months (~30 days/m)
Autonomous or Serial Line Sync Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 2.33 sec + (0.87 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.02 | 1 | 0.0466 | 20 | 0.968 | 5 | 774.7934 | 25.8264 | 371900.8264 | 23.23%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 2.33 | (sec) sampling duration | 3 | (min) sampling interval
|  |  | 180 | (sec) sampling interval
(Interval =x)
Polled Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 1.22 sec + (0.87 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.039 | 1 | 0.04758 | 6 | 0.32148 | 5 | 2332.9601 | 77.7653 | 335946.2486 | 7.716%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 1.22 | (sec) sampling duration | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  | Capacity (Kbytes) | % used
1e-05 | 3600 | 0.036 |  |  | T&C | 5 |  | 2048 | Autonom. | 0
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 37.97% | 4
[5 empty] | (no) Druck | 0 |  | Samples | Polled
using Power endurance examples SBE 37SM manual v22 |  |  |  |  | Per Sample | 9 |  | 227555.5556 | 11.39% | 1024
2048
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SM CT extpower
SBE 37SM Power Budget Calculations |  | CT, no pressure, external power
(enter variables into green boxes) |  | (click here to -> Intro) |  |  |  | Target duration = | 12 | months (~30 days/m)
Autonomous or Serial Line Sync Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 2.33 sec + (0.87 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.035 | 1 | 0.08155 | 20 | 1.667 | 5 | 449.91 | 14.997 | 215956.8086 | 80.02%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 2.33 | (sec) sampling duration | 3 | (min) sampling interval
|  |  | 180 | (sec) sampling interval
(Interval =x)
Polled Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 1.22 sec + (0.87 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.035 | 1 | 0.0427 | 20 | 0.89 | 5 | 842.6966 | 28.0899 | 404494.382 | 42.72%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 1.22 | (sec) sampling duration | 3 | (min) sampling interval
|  |  | 180 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  | Capacity (Kbytes) | % used
1e-05 | 3600 | 0.036 |  |  | T&C | 5 |  | 2048 | Autonom. | 0
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 75.94% | 4
[5 empty] | (no) Druck | 0 |  | Samples | Polled
using Power endurance examples SBE 37SM manual v22 |  |  |  |  | Per Sample | 9 |  | 227555.5556 | 75.94% | 1024
2048
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SM CTD intpower
SBE 37SM Power Budget Calculations |  | CT with pressure, internal power
(enter variables into green boxes) |  | (click here to -> Intro) |  |  |  | Target duration = | 6 | months (~30 days/m)
Autonomous or Serial Line Sync Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 2.83 sec + (1.27 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.02 | 1 | 0.0566 | 6 | 0.3756 | 5 | 1996.8051 | 66.5602 | 287539.9361 | 9.014%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 2.83 | (sec) sampling duration | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Polled Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 1.72 sec + (1.27 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.039 | 1 | 0.06708 | 6 | 0.43848 | 5 | 1710.4543 | 57.0151 | 246305.4187 | 10.52%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 1.72 | (sec) sampling duration | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  | Capacity (Kbytes) | % used
1e-05 | 3600 | 0.036 |  |  | T&C | 5 |  | 2048 | Autonom. | 0
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 13.92% | 4
[5 empty] | Druck | 2 |  | Samples | Polled
using Power endurance examples SBE 37SM manual v22 |  |  |  |  | Per Sample | 11 |  | 186181.8182 | 13.92% | 1024
2048
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SM CTD extpower
SBE 37SM Power Budget Calculations |  | CT with pressure, external power
(enter variables into green boxes) |  | (click here to -> Intro) |  |  |  | Target duration = | 6 | months (~30 days/m)
Autonomous or Serial Line Sync Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 2.83 sec + (1.27 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.035 | 1 | 0.09905 | 20 | 2.017 | 5 | 371.8394 | 12.3946 | 178482.8954 | 48.41%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 2.83 | (sec) sampling duration | 3 | (min) sampling interval
|  |  | 180 | (sec) sampling interval
(Interval =x)
Polled Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 1.72 sec + (1.27 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.035 | 1 | 0.0602 | 20 | 1.24 | 5 | 604.8387 | 20.1613 | 290322.5806 | 29.76%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 1.72 | (sec) sampling duration | 3 | (min) sampling interval
|  |  | 180 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  | Capacity (Kbytes) | % used
1e-05 | 3600 | 0.036 |  |  | T&C | 5 |  | 1024 | Autonom. | 0
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 0 | (0 or 4) | Available | 59.06% | 4
[5 empty] | (no) Druck | 2 |  | Samples | Polled
using Power endurance examples SBE 37SM manual v22 |  |  |  |  | Per Sample | 7 |  | 146285.7143 | 59.06% | 1024
2048
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SM CT RS-485
SBE 37SM Power Budget Calculations |  | CT, no pressure, RS-485
(enter variables into green boxes) |  | (click here to -> Intro)
Number of Instruments on line
| 10 | 0.0055 | 1 |  |  | Target duration = | 12 | months (~30 days/m)
|  | amp-sec per hr | requests/hr
[5 empty] | Battery | deployment |  | # samples | % used
Autonomous or Serial Line Sync Sampling |  |  |  |  | capacity (AH) | days or | months
0.035 | 2.3 | 0.0805 | 6 | 0.5965 | 5 | 1257.3345 | 41.9111 | 181056.1609 | 28.63%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
[5 empty] | Battery | deployment |  | # samples | % used
Polled Sampling |  |  |  |  | capacity (AH) | days or | months
0.035 | 1.2 | 0.042 | 1 | 0.15 | 5 | 5000 | 166.6667 | 120000 | 7.2%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 60 | (min) sampling interval
|  |  | 3600 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  | Capacity (Kbytes) | % used | 0
3e-05 | 3600 | 0.108 |  |  | T&C | 5 |  | 2048 | Autonom. | 4
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 22.78%
[5 empty] | (no) Druck | 0 |  | Samples | Polled | 1024
using Power endurance examples SBE 37SM manual v22 |  |  |  |  | Per Sample | 9 |  | 227555.5556 | 3.797% | 2048
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SM CTD RS-485
SBE 37SM Power Budget Calculations |  | CT with pressure, RS-485
(enter variables into green boxes) |  | (click here to -> Intro)
Number of Instruments on line
| 10 | 0.0055 | 1 |  |  | Target duration = | 12 | months (~30 days/m)
|  | amp-sec per hr | requests/hr
[5 empty] | Battery | deployment |  | # samples | % used
Autonomous or Serial Line Sync Sampling |  |  |  |  | capacity (AH) | days or | months
0.035 | 2.8 | 0.098 | 6 | 0.7015 | 5 | 1069.1376 | 35.6379 | 153955.809 | 33.67%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
[5 empty] | Battery | deployment |  | # samples | % used
Polled Sampling |  |  |  |  | capacity (AH) | days or | months
0.035 | 1.6 | 0.056 | 1 | 0.164 | 5 | 4573.1707 | 152.439 | 109756.0976 | 7.872%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 60 | (min) sampling interval
|  |  | 3600 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  | Capacity (Kbytes) | % used | 0
3e-05 | 3600 | 0.108 |  |  | T&C | 5 |  | 2048 | Autonom. | 4
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 27.84%
[5 empty] | (no) Druck | 2 |  | Samples | Polled | 1024
using Power endurance examples SBE 37SM manual v22 |  |  |  |  | Per Sample | 11 |  | 186181.8182 | 4.641% | 2048
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SMP CT intpower
SBE 37SMP Power Calculations |  | CT, no pressure, internal power
(enter variables into green boxes) |  | (click here to -> Intro) |  |  |  | Target duration = | 6 | months (~30 days/m)
Autonomous or Serial Line Sync Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 2.83 sec + (1.87 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.02 | 1 | 0.0566 | 6 | 1.1556 | 5 | 649.0135 | 21.6338 | 93457.9439 | 27.73%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 2.83 | (sec) sampling duration | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Polled Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 2.24 sec + (1.87 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.039 | 1 | 0.08736 | 6 | 1.3402 | 5 | 559.6347 | 18.6545 | 80587.3926 | 32.16%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 2.24 | (sec) sampling duration | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  |  | % used
1e-05 | 3600 | 0.036 |  |  | T&C | 5 |  |  | Autonom. | 0
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 11.39% | 4
Pump current consumption for 0.5 sec |  |  |  |  | (no) Druck | 0 |  | Samples | Polled
0.26 | 0.5 | 0.13 |  |  | Per Sample | 9 |  | 227555.5556 | 11.39%
amps | sampling duration (s) | amp-sec per hr used
using Power endurance examples SBE 37SMP manual v3
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SMP CT extpower
SBE 37SMP Power Calculations |  | CT, no pressure, external power
(enter variables into green boxes) |  | (click here to -> Intro) |  |  |  | Target duration = | 6 | months (~30 days/m)
Autonomous or Serial Line Sync Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 2.83 sec + (1.87 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.035 | 1 | 0.09905 | 6 | 1.4103 | 5 | 531.8017 | 17.7267 | 76579.4512 | 33.85%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 2.83 | (sec) sampling duration | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Polled Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 2.24 sec + (1.87 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.035 | 1 | 0.0784 | 6 | 1.2864 | 5 | 583.0224 | 19.4341 | 83955.2239 | 30.87%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 2.24 | (sec) sampling duration | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  |  | % used
1e-05 | 3600 | 0.036 |  |  | T&C | 5 |  |  | Autonom. | 0
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 11.39% | 4
Pump current consumption for 0.5 sec |  |  |  |  | (no) Druck | 0 |  | Samples | Polled
0.26 | 0.5 | 0.13 |  |  | Per Sample | 9 |  | 227555.5556 | 11.39%
amps | sampling duration (s) | amp-sec per hr used
using Power endurance examples SBE 37SMP manual v3
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SMP CTD intpower
SBE 37SMP Power Calculations |  | CT with pressure, internal power
(enter variables into green boxes) |  | (click here to -> Intro) |  |  |  | Target duration = | 6 | months (~30 days/m)
Autonomous or Serial Line Sync Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 3.33 sec + (2.27 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.02 | 1 | 0.0666 | 6 | 1.2156 | 5 | 616.9793 | 20.566 | 88845.0148 | 29.17%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 3.33 | (sec) sampling duration | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Polled Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 2.74 sec + (2.27 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.039 | 1 | 0.10686 | 6 | 1.4572 | 5 | 514.6998 | 17.1567 | 74116.7751 | 34.97%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 2.74 | (sec) sampling duration | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  |  | % used
1e-05 | 3600 | 0.036 |  |  | T&C | 5 |  |  | Autonom. | 0
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 13.92% | 4
Pump current consumption for 0.5 sec |  |  |  |  | (no) Druck | 2 |  | Samples | Polled
0.26 | 0.5 | 0.13 |  |  | Per Sample | 11 |  | 186181.8182 | 13.92%
amps | sampling duration (s) | amp-sec per hr used
using Power endurance examples SBE 37SMP manual v3
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SMP CTD extpower
SBE 37SMP Power Calculations |  | CT with pressure, external power
(enter variables into green boxes) |  | (click here to -> Intro) |  |  |  | Target duration = | 6 | months (~30 days/m)
Autonomous or Serial Line Sync Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 3.33 sec + (2.27 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.035 | 1 | 0.11655 | 6 | 1.5153 | 5 | 494.9515 | 16.4984 | 71273.0152 | 36.37%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 3.33 | (sec) sampling duration | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Polled Sampling |  |  |  |  | Battery | deployment |  | # samples | % used
= 2.74 sec + (2.27 sec * (NAVG-1)) |  |  |  |  | capacity (AH) | days or | months
0.035 | 1 | 0.0959 | 6 | 1.3914 | 5 | 539.0254 | 17.9675 | 77619.6636 | 33.39%
amps | NAVG (default = 1) | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
range = 1 to 1000
| 2.74 | (sec) sampling duration | 10 | (min) sampling interval
|  |  | 600 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  |  | % used
1e-05 | 3600 | 0.036 |  |  | T&C | 5 |  |  | Autonom. | 0
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 13.92% | 4
Pump current consumption for 0.5 sec |  |  |  |  | (no) Druck | 2 |  | Samples | Polled
0.26 | 0.5 | 0.13 |  |  | Per Sample | 11 |  | 186181.8182 | 13.92%
amps | sampling duration (s) | amp-sec per hr used
using Power endurance examples SBE 37SMP manual v3
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SMP CT RS-485
SBE 37SMP RS-485 Power Calculations |  | CT, no pressure | (enter variables into green boxes)
(click here to -> Intro)
Number of Instruments on line
| 10 | 0.025 | 1 |  |  | Target duration = | 12 | months (~30 days/m)
|  | amp-sec per hr | requests/hr
[5 empty] | Battery | deployment |  | # samples | % used
Autonomous or Serial Line Sync Sampling |  |  |  |  | capacity (AH) | days or | months
0.035 | 3.3 | 0.1155 | 12 | 3.079 | 5 | 243.5856 | 8.1195 | 70152.647 | 147.8%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 5 | (min) sampling interval
|  |  | 300 | (sec) sampling interval
(Interval =x)
[5 empty] | Battery | deployment |  | # samples | % used
Polled Sampling |  |  |  |  | capacity (AH) | days or | months
0.035 | 2.2 | 0.077 | 1 | 0.47 | 5 | 1595.7447 | 53.1915 | 38297.8723 | 22.56%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 60 | (min) sampling interval
|  |  | 3600 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  |  | % used | 0
3e-05 | 3600 | 0.108 |  |  | T&C | 5 |  |  | Autonom. | 4
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 45.56%
Pump current consumption for 0.5 sec |  |  |  |  | (no) Druck | 0 |  | Samples | Polled | 1024
0.26 | 0.5 | 0.13 |  |  | Per Sample | 9 |  | 227555.5556 | 3.797% | 2048
amps | sampling duration (s) | amp-sec per hr
Using Power endurance examples SBE 37IMP manual v3, w/ 1.75 seconds added to polled sample
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 37SMP CTD RS-485
SBE 37SMP RS-485 Power Calculations |  | CT with pressure | (enter variables into green boxes)
(click here to -> Intro)
Number of Instruments on line
| 10 | 0.025 | 1 |  |  | Target duration = | 12 | months (~30 days/m)
|  | amp-sec per hr | requests/hr
[5 empty] | Battery | deployment |  | # samples | % used
Autonomous or Serial Line Sync Sampling |  |  |  |  | capacity (AH) | days or | months
0.035 | 3.8 | 0.133 | 30 | 8.023 | 5 | 93.4812 | 3.116 | 67306.4938 | 385.1%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 2 | (min) sampling interval
|  |  | 120 | (sec) sampling interval
(Interval =x)
[5 empty] | Battery | deployment |  | # samples | % used
Polled Sampling |  |  |  |  | capacity (AH) | days or | months
0.035 | 2.6 | 0.091 | 10 | 2.473 | 5 | 303.2754 | 10.1092 | 72786.0898 | 118.7%
amps | seconds of sampling | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
|  |  | 6 | (min) sampling interval
|  |  | 360 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  |  |  | Memory |  |  |  | % used | 0
3e-05 | 3600 | 0.108 |  |  | T&C | 5 |  |  | Autonom. | 4
amps | sec/hr | amp-sec per hr |  |  | Time Stamp | 4 | (0 or 4) | Available | 139.2%
Pump current consumption for 0.5 sec |  |  |  |  | (no) Druck | 2 |  | Samples | Polled | 1024
0.26 | 0.5 | 0.13 |  |  | Per Sample | 11 |  | 186181.8182 | 46.41% | 2048
amps | sampling duration (s) | amp-sec per hr
Using Power endurance examples SBE 37IMP manual v3, w/ 1.75 seconds added to polled sample
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 39 interval
SBE 39 Interval Operation Power Budget Calculation |  |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) | (alkaline only) |  | Note: enter 20 if using lithium batteries
polled sampling; interval, autonomous |  |  |  | Target duration = | 12 | months (~30 days/m)
sampling; or serial line sync sampling
|  |  | Battery | deployment |  | # samples | % used
No Pressure sensor |  |  | capacity (AH) | days or | months
0.018 | 10 | 0.216 | 1.1 | 763.8889 | 25.463 | 183333.3333 | 47.13%
amp-sec per sample | # of samples/hr | amp-sec/hr used | (1.1 for Lithium)
acquisition current | 6 | (min) sampling interval | (0.3 for Alkaline for warm water)
| 360 | (sec) sampling interval
(Interval =x)
|  |  | Battery | deployment |  | # samples | % used
Pressure sensor installed |  |  | capacity (AH) | days or | months
0.023 | 12 | 0.312 | 1.1 | 528.8462 | 17.6282 | 152307.6923 | 68.07%
amp-sec per sample | # of samples/hr | amp-sec/hr used | (1.1 for Lithium)
acquisition current | 5 | (min) sampling interval | (0.3 for Alkaline for warm water)
| 300 | (sec) sampling interval
(Interval =x)
Quiescent consumption in AH |  |  | Memory |  |  | Capacity (Kbytes) | % used | 0
1e-05 | 3600 | 0.036 |  |  |  | 2048 | no Druck | 2
amps | sec/hr | amp-sec per hr | Temp/Time | 7 |  | Available | 37.97%
|  |  | Pressure | 2 |  | Samples | with Druck | 2048
|  |  | Per Sample | 9 |  | 227555.5556 | 45.56% | 4096
using Power endurance examples SBE 39 manual v16, p10
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 39 continuous
SBE 39 Power Budget Calculations |  | Avg. Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
(click here to -> Intro)
continuous sampling | battery |  |  | deployment capability
| capacity (AH) |  |  | in hrs | or in days
0.015 | 1.1 |  |  | 73.3333 | 3.0556
amp-sec per sample | (1.1 for Lithium)
acquisition current | (0.3 for Alkaline for warm water) |  |  | duration | Number of
|  |  |  | (seconds) | samples | % of Memory
|  |  | Temp only, Real time = N | 0.8 | 330000 | 112.8%
|  |  | Temp only, Real time = Y | 1 | 264000 | 90.23%
|  |  | T&P, Real time = N | 1.25 | 211200 | 72.19%
|  |  | T&P, Real time = Y | 1.5 | 176000 | 60.16%
2048
|  | Memory |  |  | Capacity (kbyte) |  | 4096
|  | Temp/Time | 7 | bytes | 2048
|  | Pressure | 0 | bytes | available samples |  | 0
|  | Per Sample | 7 | bytes | 292571.4286 |  | 2
using Power endurance examples SBE 39 manual v16, p10 and INTERVAL = X (= 0) on page 24
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 44 internal power
SBE 44 Power Budget Calculations |  | (enter variables into green boxes)
(click here to -> Intro)
Communication
[5 empty] | capacity (AH) | hrs | days | yrs
0.01 | 210 | 2.1 | 1 | 2.46 | 5 | 7317.0732 | 304.878 | 0.83528 | 7317.0732
amps | communication time | amp-sec per hr | # of samples/hr | amp-sec/hr used | (7.2AH is nominal, 5AH is conservative)
| per wake up [sec.] |  | 60 | (minute interval)
Power to External Device (optional, may not apply)
0 | 5 | 0 | 1
amps | On-time | amp-sec per hr | # of samples/hr
|  |  | 60 | (minute interval)
Quiescent consumption in AH
0.0001 | 3600 | 0.36
amps | sec/hr | amp-sec per hr
Using specification from SBE 44 manual v12, p10
Notes:
Deployments of batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years

## 16 CT no pump
SBE 16 Power Budget Calculations |  | CT, no pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
[6 empty] | Target duration = | 6 | months (~30 days/m)
Sampling duration is by default 5 seconds |  |  |  |  | Battery | deployment |  | # samples | % used
[5 empty] | capacity (AH) | days or | months
0.13 | 5 | 0.65 | 5 | 3.43 | 12.2 | 533.5277 | 17.7843 | 64023.3236 | 33.74%
amps | sampling duration (s) | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
|  |  |  | (add quiescent&auxes) | use 10.5 amp-hours for pump or aux sensors
|  |  | 12 | (min) interval
|  |  | 720 | (sec) interval
current draw (mA)
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | Memory | bytes used |  | Capacity (Kb) | % used | 64
Aux 2 | 0 | enter a value = 45 |  |  | T&C | 4 | 4 | 128 |  | 128
Aux 3 | 0 |  |  |  | Aux 1, 2 (pair) | 3 | 0 | Available |  | 256
Aux 4 | 0 |  |  |  | Aux 3, 4 (pair) | 3 | 0 | Samples |  | 512
Sum of Auxes | 0 | mA |  |  |  | Sum | 4 | 32000 | 67.5% | 1024
2048
Quiescent consumption in AH | [9 empty] | 4096
5e-05 | 3600 | 0.18 | [7 empty] | 8192
amps | sec/hr | amp-sec per hr
using SBE 16 manual v16.10, also applies to SBE 19 in moored mode
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
Storage Capability (on-board memory) is approximate; i.e. for '1024' Kbytes, there are actually 1043 Kbytes (1048k capacity - 5k scratchpad)

## 16 CTD no pump
SBE 16 Power Budget Calculations |  | CTD, no pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
[6 empty] | Target duration = | 3 | months (~30 days/m)
Sampling duration is by default 5 seconds |  |  |  |  | Battery | deployment |  | # samples | % used
for PAINE installed pressure sensor |  |  |  |  | capacity (AH) | days or | months
0.18 | 5 | 0.9 | 10 | 9.18 | 12.2 | 199.3464 | 6.6449 | 47843.1373 | 45.15%
amps | sampling duration (s) | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
|  |  |  | (add quiescent&auxes) | use 10.5 amp-hours for pump or aux sensors
|  |  | 6 | (min) interval
|  |  | 360 | (sec) interval
Sampling duration is by default 6.7 seconds |  |  |  |  | Battery | deployment |  | # samples | % used
for DIGIQUARTZ installed pressure sensor |  |  |  |  | capacity (AH) | days or | months
0.245 | 6.7 | 1.6415 | 15 | 24.8025 | 12.2 | 73.7829 | 2.4594 | 26561.8385 | 122%
amps | sampling duration (s) | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
|  |  |  | (add quiescent&auxes) | use 10.5 amp-hours for pump or aux sensors |  |  |  |  | 64
|  |  | 4 | (min) interval | [5 empty] | 128
|  |  | 240 | (sec) interval | [5 empty] | 256
current draw (mA) | [9 empty] | 512
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, | [7 empty] | 1024
Aux 2 | 0 | enter a value = 45 |  |  | Memory | bytes used |  | Capacity (Kb) | % used | 2048
Aux 3 | 0 |  |  |  | T&C | 4 | 4 | 64 |  | 4096
Aux 4 | 0 |  |  |  | w/ Paine | 2 | 2 |  |  | 8192
Sum of Auxes | 0 | mA |  |  | w/ Digiquartz | 5 | 0 |  | PAINE
[5 empty] | Aux 1, 2 (pair) | 3 | 0 | ~Available | 202.5%
Quiescent consumption in AH |  |  |  |  | Aux 3, 4 (pair) | 3 | 0 | Samples | DIGIQTZ
5e-05 | 3600 | 0.18 |  |  | Sum |  | 6 | 10666.6667 | 303.8%
amps | sec/hr | amp-sec per hr
using SBE 16 manual v16.10, also applies to SBE 19 in moored mode
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
Storage Capability (on-board memory) is approximate; i.e. for '1024' Kbytes, there are actually 1043 Kbytes (1048k capacity - 5k scratchpad)

## 16 CT pumped
SBE 16 Power Budget Calculations |  | CT with pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
[6 empty] | Target duration = | 6 | months (~30 days/m)
Sampling duration is by default 5 seconds |  |  |  |  | Battery | deployment |  | # samples | % used
[5 empty] | capacity (AH) | days or | months
0.13 | 5 | 2.0412 | 5 | 10.386 | 10.5 | 151.6464 | 5.0549 | 18197.5737 | 118.7%
amps | sampling duration (s) | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
0.296 | (added pump current) |  |  | (add quiescent&auxes) | use 10.5 amp-hours for pump or aux sensors
|  |  | 12 | (min) interval
|  |  | 720 | (sec) interval
current draw (mA)
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | Memory | bytes used |  | Capacity (Kb) | % used | 64
Aux 2 | 0 | enter a value = 45 |  |  | T&C | 4 | 4 | 256 |  | 128
Aux 3 | 0 |  |  |  | Aux 1, 2 (pair) | 3 | 0 | Available |  | 256
Aux 4 | 0 |  |  |  | Aux 3, 4 (pair) | 3 | 0 | Samples |  | 512
Sum of Auxes | 0 | mA |  |  |  | Sum | 4 | 64000 | 33.75% | 1024
2048
Quiescent consumption in AH | [9 empty] | 4096
5e-05 | 3600 | 0.18 | [7 empty] | 8192
amps | sec/hr | amp-sec per hr
using SBE 16 manual v16.10, also applies to SBE 19 in moored mode
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
Storage Capability (on-board memory) is approximate; i.e. for '1024' Kbytes, there are actually 1043 Kbytes (1048k capacity - 5k scratchpad)

## 16 CTD pumped
SBE 16 Power Budget Calculations |  | CTD with pump |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
[6 empty] | Target duration = | 3 | months (~30 days/m)
Sampling duration is by default 5 seconds |  |  |  |  | Battery | deployment |  | # samples | % used
for PAINE installed pressure sensor |  |  |  |  | capacity (AH) | days or | months
0.18 | 5 | 2.2912 | 10 | 23.092 | 10.5 | 68.2054 | 2.2735 | 16369.3054 | 132%
amps | sampling duration (s) | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
0.296 | (added pump current) |  |  | (add quiescent&auxes) | use 10.5 amp-hours for pump or aux sensors
|  |  | 6 | (min) interval
|  |  | 360 | (sec) interval
Sampling duration is by default 6.7 seconds |  |  |  |  | Battery | deployment |  | # samples | % used
for DIGIQUARTZ installed pressure sensor |  |  |  |  | capacity (AH) | days or | months
0.245 | 6.7 | 3.5063 | 1 | 3.6863 | 10.5 | 427.2577 | 14.2419 | 10254.1844 | 21.06%
amps | sampling duration (s) | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
0.296 | (added pump current) |  |  | (add quiescent&auxes) | use 10.5 amp-hours for pump or aux sensors |  |  |  |  | 64
|  |  | 60 | (min) interval | [5 empty] | 128
|  |  | 3600 | (sec) interval | [5 empty] | 256
current draw (mA) | [9 empty] | 512
Aux 1 | 0 | e.g. if the WETStar @ 10V draws ~45mA, | [7 empty] | 1024
Aux 2 | 0 | enter a value = 45 |  |  | Memory | bytes used |  | Capacity (Kb) | % used | 2048
Aux 3 | 0 |  |  |  | T&C | 4 | 4 | 64 |  | 4096
Aux 4 | 0 |  |  |  | w/ Paine | 2 | 2 |  |  | 8192
Sum of Auxes | 0 | mA |  |  | w/ Digiquartz | 5 | 0 |  | PAINE
[5 empty] | Aux 1, 2 (pair) | 3 | 0 | ~Available | 202.5%
Quiescent consumption in AH |  |  |  |  | Aux 3, 4 (pair) | 3 | 0 | Samples | DIGIQTZ
5e-05 | 3600 | 0.18 |  |  | Sum |  | 6 | 10666.6667 | 20.25%
amps | sec/hr | amp-sec per hr
using SBE 16 manual v16.10, also applies to SBE 19 in moored mode
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
Storage Capability (on-board memory) is approximate; i.e. for '1024' Kbytes, there are actually 1043 Kbytes (1048k capacity - 5k scratchpad)

## 16 CT latched DO
SBE 16 Power Budget Calculations |  | CT with pump & DO |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
[6 empty] | Target duration = | 2 | months (~30 days/m)
Sampling duration is by default 20 seconds |  |  |  |  | Battery | deployment |  | # samples | % used
[5 empty] | capacity (AH) | days or | months
0.13 | 20 | 9.3912 | 3 | 28.3536 | 10.5 | 55.5485 | 1.8516 | 3999.4921 | 108%
amps | sampling duration (s) | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
0.296 | (added pump current) |  |  | (add quiescent&auxes) | use 10.5 amp-hours for pump or aux sensors
|  |  | 20 | (min) interval
|  |  | 1200 | (sec) interval
current draw (mA)
Aux 1&2 DO | 8 | mA |  |  | Memory | bytes used |  | Capacity (Kb) | % used | 64
[5 empty] | T&C | 4 | 4 | 64 |  | 128
Aux 3 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | Aux 1, 2 (pair) | 3 | 0 | Available |  | 256
Aux 4 | 0 | enter a value = 45 |  |  | Aux 3, 4 (pair) | 3 | 0 | Samples |  | 512
Sum of Auxes 3&4 | 0 | mA |  |  |  | Sum | 4 | 16000 | 27% | 1024
2048
Quiescent consumption in AH | [9 empty] | 4096
5e-05 | 3600 | 0.18 | [7 empty] | 8192
amps | sec/hr | amp-sec per hr
using SBE 16 manual v16.10, also applies to SBE 19 in moored mode
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
Storage Capability (on-board memory) is approximate; i.e. for '1024' Kbytes, there are actually 1043 Kbytes (1048k capacity - 5k scratchpad)

## 16 CTD latched DO
SBE 16 Power Budget Calculations |  | CTD with pump & DO |  | Average Water Temp °C | 20 | ( 0°C <value< 20°C, enter 20 if >20)
(enter variables into green boxes) |  | (click here to -> Intro) |  | (alkaline only) |  | Note: enter 20 if using lithium batteries
[6 empty] | Target duration = | 3 | months (~30 days/m)
Sampling duration is by default 20 seconds |  |  |  |  | Battery | deployment |  | # samples | % used
for PAINE installed pressure sensor |  |  |  |  | capacity (AH) | days or | months
0.18 | 20 | 10.3912 | 1 | 10.5712 | 10.5 | 148.9897 | 4.9663 | 3575.753 | 60.41%
amps | sampling duration (s) | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
0.296 | (added pump current) |  |  | (add quiescent&auxes) | use 10.5 amp-hours for pump or aux sensors
Sampling duration is by default 20 seconds |  |  |  |  | Battery | deployment |  | # samples | % used
for DIGIQUARTZ installed pressure sensor |  |  |  |  | capacity (AH) | days or | months
0.245 | 20 | 11.6616 | 1 | 11.8416 | 10.5 | 133.0057 | 4.4335 | 3192.1362 | 67.67%
amps | sampling duration (s) | amp-sec per hr | # of samples/hr | amp-sec/hr used | use 12.2 amp-hours for no pump or auxiliary sensors
0.296 | (added pump current) |  |  | (add quiescent&auxes) | use 10.5 amp-hours for pump or aux sensors |  |  |  |  | 64
128
256
current draw (mA) | [9 empty] | 512
Aux 1&2 DO | 8 | mA | [7 empty] | 1024
[5 empty] | Memory | bytes used |  | Capacity (Kb) | % used | 2048
Aux 3 | 0 | e.g. if the WETStar @ 10V draws ~45mA, |  |  | T&C | 4 | 4 | 256 |  | 4096
Aux 4 | 0 | enter a value = 45 |  |  | w/ Paine | 2 | 2 |  |  | 8192
Sum of Auxes 3&4 | 0 | mA |  |  | w/ Digiquartz | 5 | 0 |  | PAINE
[5 empty] | Aux 1, 2 (pair) | 3 | 0 | ~Available | 5.062%
Quiescent consumption in AH |  |  |  |  | Aux 3, 4 (pair) | 3 | 0 | Samples | DIGIQTZ
5e-05 | 3600 | 0.18 |  |  | Sum |  | 6 | 42666.6667 | 5.062%
amps | sec/hr | amp-sec per hr
using SBE 16 manual v16.10
Note: Storage Capability (on-board memory) is approximate; i.e. for '1024' Kbytes, there are actually 1043 Kbytes (1048k capacity - 5k scratchpad)
Notes:
Deployments of alkaline batteries longer than 2 years are not advised, regardless of calculated deployment life
Shelf Life for Lithium battteries is ~10 years
Also applies to SBE 19 in moored mode